
Modelling of the Neutron Production in a Mixed Beam DT Neutron Generator
Aljaž Čufara, Paola Batistonib, Zamir Ghanic, Luca Giacomellid, Igor Lengara, Stefano Loretib,
Alberto Miloccoe, Sergey Popovichevc, Mario Pillonb, Davide Rigamontid,e, Marica Rebaid,e,
Marco Tardocchid, Luka Snoja,f, JET contributors1
EUROfusion consortium, Culham Science Centre, Abington, United Kingdom

aJožef Stefan Institute, Ljubljana, Slovenia

bENEA, I-00044 Frascati, Rome, Italy

cCCFE, Abingdon, Oxon, OX14 3DB, United Kingdom

dIstituto di Fisica del Plasma “P. Caldirola”, CNR, Milano, Italy

eDipartimento di Fisica “G. Occhialini”, Università degli Studi di Milano-Bicocca, Milano, Italy
fFaculty of mathematics and physics, University of Ljubljana, Slovenia
Compact DT neutron generators based on accelerators are often built on the principle of a mixed beam operation, meaning that deuterium (D) and tritium (T) are both present in the ion beam and in the target. Moreover, the beam consists of a mixture of ions and ionized molecules (D, T ions, and ionized D-D, T-T and D-T molecules) so the relevant source components come from T(d, n), D(t, n), D(d, n) and T(t, 2n) reactions at different ion energies.
The method for assessing the relative amplitudes of different source components (DD, DT, TT) is presented. The assessment relies on the measurement of the neutron spectrum of different DT components (T(d, n) and D(t, n) at different energies) using a high resolution neutron spectrometer, e.g. a diamond detector, fusion reaction cross-sections, and simulations of neutron generation in the target. Through this process a complete description of the neutron source properties of the mixed beam neutron generator can be made and a neutron source description card, in a format suitable for Monte Carlo code MCNP, produced.
Keywords: DT generator, mixed beam, neutron emission, ENEA-JSI subroutine, MCUNED
1. Introduction

Compact DT neutron generators suitable for use in calibrations of neutron yield detectors to DT neutrons can be based on mixed beam operation in which mixed beam and target, i.e. D and T are present in both target and the beam, are used to prolong the operational life of the generator by continuously implanting D and T in the target during the generator’s operation. Due to the lack of an analyzing magnet, the ion beam consists of both ions and ionized molecules.
The nuclei interacting in the target can thus originate from ions (D+, T+) or ionized molecules (D-D+, D-T+, T-T+) accelerated to energy defined by the acceleration system. Heavier (triatomic) molecules in the beam are neglected. The nuclei originating from ions receive full acceleration energy (Eion) while the nuclei from ionized molecules only receive a part of the energy based on their mass relative to the total mass of the molecule. This results in complex neutron emission properties with six distinct DT, three DD, and three TT source components. However, due to low relative intensity of DD and TT source components, only DD and TT neutrons generated by D+ and T+ beam ions can be considered in the source definition.  In this case, the source description consists of eight components, i.e. six DT, one DD, and one TT source component.
The results and experience presented in the present paper are based on the process used in the characterization of the ING-17 14 MeV neutron generator, provided by VNIIA (Russia), which was used in the latest calibration of neutron yield monitors at JET  (intensity ≈ 2x108 n/s) [1]. According to the supplier, both the generator beam and target contain about 50% D and 50% T.  
2. Relative intensities of different neutron source components
2.1 Fusion reaction cross-sections
The cross-sections for fusion reactions of interest are in Fig. 1. At energies relevant for typical compact neutron generators (up to 110 keV [2]), the cross-sections for DD and TT reactions are significantly lower than for DT reaction. This limits the importance of the DD and TT components in typical generators with approximately equal split between D and T in both the target and ion beam [2].
[image: image1.png]Cross-section (barn)

an-4 - = -D(t’ n) ]
o 4p: D(d, n)
- T(t, 2n) ]

0 50 100 150 200 250 300 350 400
lon energy (keV)




Fig. 1. Cross-sections for DT, DD and TT fusion reactions in the laboratory coordinate system [3].
2.2 DT components
In the mixed beam neutron generator based on the accelerator which accelerates positive ions and ionized molecules to energy Eion, in total six DT source components are expected. These components originate from fusion reactions:

· T(d, n) at deuterium ion energies Eion, 1/2 Eion and 2/5 Eion , from D+, D-D+ and D-T+, respectively, and

· D(t, n) at tritium ion energies Eion, 1/2 Eion and 3/5 Eion from T+, T-T+ and D-T+, respectively.
The spectra of each of the source components were calculated using the ENEA-JSI source subroutine [4] added to the standard MCNP6 [5]; however, other codes were also considered [6].

To perform the neutron source simulations, first the ion acceleration voltage needs to be known. This voltage does not necessarily correspond to the voltage indicated by the neutron generator control unit as there can be a significant drop in voltage between the high voltage generator and the ion acceleration stage of the neutron generator. However, as the acceleration voltage determines the energies of ions in T(d, n) and D(t, n) fusion reactions, its determination is crucial for accurate simulation of the source components.
Once Eion is determined, all six DT source components are simulated separately to get neutron emission spectra. Moreover, it can be assumed that the source components from D-T+ molecules (T(d, n) at 2/5 and D(t, n) at 3/5 of Eion) are 100 % correlated. As suggested by the supplier and the literature [2], equal amounts of D and T in the target (50%D, 50%T) were assumed. The neutron yields of different DT source components are shown in Fig. 2 and are given in Table 1 relative to the T(d, n) component resulting from D+ ions at 73 keV. Relative neutron yields and 100 % correlation of T(d, n) at 2/5 Eion and D(t, n) at 3/5 Eion are then be used to decrease the number of fitting parameters to five.
[image: image2.png]Neutron yield per ion (n/ion)

’ —=—T(d,n) ]
--e--D(t, n) | j
/ —+—D(d, n) ]
T(t, 2n) ]

20 40 60 80 100

120




Fig. 2. Neutron yields, i.e. probabilities for neutron production per ion, of relevant reactions as a function of Eion.

Table 1. Neutron yields of source components, i.e. probabilities for neutron production per ion, relative to the T(d, n) component resulting from D+ ions at 73 keV.
	Fusion reaction
	Ion energy (keV)
	Relative neutron yield (%)

	T(d, n)
	73.0
	100

	T(d, n)
	36.5
	7.24

	T(d, n)
	29.2
	2.66

	D(t, n)
	73.0
	35.5

	D(t, n)
	36.5
	1.71

	D(t, n)
	43.8
	4.11


In the characterization of the neutron generator used for JET neutron monitor calibration, the ion acceleration voltage of 73 keV was obtained by comparing the calculated neutron spectrum with the spectrum obtained by a high resolution spectroscopic measurement using a diamond detector located directly in front of the neutron generator (Fig. 3) [7], [8]. To perform the fitting, the detector’s response function was used. This detector position is used due to the largest differences in energies of the peaks being at angles of 0° and 180° relative to the direction of the ion beam. The measured neutron spectrum could be reproduced by a linear combination of the six DT components described above. Generally the highest energy neutrons correspond to D(t, n) at T+ ion energies Eion. More details on the fitting process can be found in [7], and results relevant for the compact DT neutron generator used in the latest calibration of the neutron yield detectors in JET [1] in Fig. 4 and Table 2.
[image: image3.png]neutron generator tube

ion beam

1984e)

JarawoJldads





Fig. 3. Sketch of the neutron spectrum measurement process.
DD and TT neutron source components cannot be determined the same way as DT components due to the complex response function of the diamond detector [8] at energies relevant for DD leading to significantly decreased energy resolution, and due to wide emission spectrum of the TT component making it difficult to discern from the slowed down DT neutrons.
[image: image4.png]Counts / bin

1200 T T T T T

1000 L ® Counts NG1
—— Numerical Fit
800
600
400

200

0 L L L
8.2 8.4 8.6 8.8 9 9.2
Edep [MeV]





Fig. 4. Measured spectrum of α particles resulting from C(n, α) in diamond, and best fit of components from simulations [7].
Table 2. Relative intensities of DT neutron source components.
	Fusion reaction
	Ion energy (keV)
	Relative intensity (%)

	T(d, n)
	73.0
	7.32

	T(d, n)
	36.5
	15.14

	T(d, n)
	29.2
	24.95

	D(t, n)
	73.0
	11.06

	D(t, n)
	36.5
	2.93

	D(t, n)
	43.8
	38.61


2.3 Beam composition
The measured neutron spectrum could be reproduced by a linear combination of the DT components described in the previous section. Using this linear combination, the ion beam components can be determined by the use of the relative neutron yields from neutron source simulations using the MCUNED code [10] (Table 1) and the source component intensities from Table 2. For the target composition, typically equal concentrations of D and T, and the ratio of hydrogen to titanium atoms of up to 2 to 1 has been assumed [2]. The ion beam composition determined through this process is presented in Table 3.
Table 3. Ion beam composition determined from fitted neutron source component intensities and assumed target composition.

	Ion beam
component
	Relative abundance in the ion beam (%)

	D+
	0.63

	T+
	2.66

	D-D+
	8.95

	T-T+
	7.32

	D-T+
	80.44


In principle the neutron spectrum can change throughout the operational life of the generator due to changes in its target and beam composition, e.g. as a result of T decay, and D or T building-up or escaping from the target. However, the neutron spectrum measurements at different times in the characterization process, i.e. spectra after different operational histories of the generator, showed no major changes in the neutron emission spectrum of the generator used. 

2.4 Relative importance of DD and TT components
A first approximation for the upper limit of the relative intensities of the DD and TT source components is based entirely on the fusion reaction cross-sections. For both ion beam and target containing equal amounts of D and T, the DD and TT contribution to the total neutron production can be calculated using Equations 1 and 2. Contributions calculated based on the fusion cross-sections from Fig. 1 are in Fig. 5. σDT, σTD, σDD, and σTD are the cross-sections for T(d, n), D(t, n), D(d, n), and T(t, 2n), respectively.

[image: image5.wmf]2

DD

DTTDDDTT

DDcontribution

s

ssss

=

+++×


(1)


[image: image6.wmf]2

2

TT

DTTDDDTT

TTcontribution

s

ssss

×

=

+++×


(2)

The average energy at which the ion from the beam interacts with D or T in the target is lower than the acceleration energy Eion, hence the maximum of the relative DD and TT contributions calculated on the energy interval [Emin, Eion] represent the upper limit of intensity of these components. Emin, the lowest energy that can realistically still give significant contribution to the neutron production, is selected depending on the Eion and the cross-sections for fusion reactions (Fig. 1). For most cases Emin between 10 and 20 keV seems like a reasonable choice.
[image: image7.png]50 T T T T T T T T T T T T

45_- ........... DD Component
4.0 TT component ]
1——DD + TT components

Relative contribution (%)

0 50 100 150 200 250 300 350
lon energy (keV)




Fig. 5. Relative contributions of DD and TT source components to the total fusion cross-section as a function of ion energy.
For typical Eion of up to 110 keV, the highest upper limit for the DD and TT source components seem to be approximately 0.4 %, while the combined intensity, i.e. the upper limit on the intensity of non DT source components, does not exceed 0.7 % of the total neutron emission in Fig. 5. This means that in case of a compact neutron generator with typical Eion (up to 110 keV) the DD and TT source components can in most cases be neglected unless 1 % uncertainty represents significant addition to the total uncertainty. However, an increase of the acceleration voltage to 300 kV, as is typically done in neutron generators with higher source intensities, can potentially results in significantly higher intensities of components other than DT. In such cases the modelling of DD and TT components is advised. A possible use for the mixed beam neutron generator with relatively high intensity, and thus possibly high acceleration voltage, will be the calibration source used in calibration of neutron detectors to DT neutrons in ITER.
Finally, in order to assess the relative importance of DD and TT components, full simulation of the neutron source was performed in a simple model where only a  thin 2 µm-target and 73 keV ion beam were present. The simulation was repeated also for the case of a 10 µm target and 300 keV ion beam to compare with the case of a neutron generator with higher intensity. Such a model was chosen to minimize neutron scattering and thus to make it easier to separate different source components in the neutron emission spectrum. In our analysis the target consisted entirely of Ti-D-T, i.e. there were two hydrogen atoms per Ti atom, and an equal split between D and T was assumed. Each ion beam component, i.e. D+ at Eion, 1/2 Eion and 2/5 Eion, and T+ at Eion, 1/2 Eion and 3/5 Eion, can be simulated separately.
Due to relatively low amount of neutron scattering in the target, the different components in the emission spectra can be separated. There can be some difficulties in discerning between TT neutrons and down-scattered DT neutrons; however, previous analyses of DT source with no DD or TT components showed that in geometry where only the thin target is filled with material, there is only 0.01 % of neutrons at energies below the DT peak [11]. In the present analysis the down-scattered DT neutrons were approximated as a constant background to DD and TT components. A more detailed and rigorous analysis was not deemed necessary due to the low number of neutrons scattered to lower energies.

[image: image8.png]Neutron spectrum (n / MeV)

— — -DT component
DD component
TT component

2 4 6 8 10 12

Neutron energy (MeV)

14





Fig. 6. The simulated neutron emission spectrum of a mixed beam neutron generator operating at 73 keV neutron generator. The colored area represents the sum of all components.
Results of the analysis are shown in Fig. 6 and Table 4. The analysis for the neutron generator with the same beam composition (ion beam composition from Table 3), i.e. approximation using the same kind of ion source, and acceleration voltage of 300 keV was performed to test the performance at higher ion energies.

Based on these results, the DT neutron source components of a compact neutron generator used in JET’s latest DT calibration represent roughly 99 % of the neutrons emitted from its target. Even lower importance of the components other than DT were found for the neutron generator built on the same principle but with 300 kV acceleration voltage. This is consistent with the fact that most of the ion beam represent D-T+ molecules (Table 3), slowing down of ions in the target and relative contributions shown in Fig. 5.

To take DD and TT components into account the relative intensities of different DT components need to be multiplied by the relative intensity of all DT components (roughly 99 % in both presented cases), and the other two components added according to their relative intensities (Table 4).
Table 4. The relative intensity of the different neutron source components for 73 keV and 300 keV acceleration energy.

	Source component
	Relative intensity (%)

	Eion
	73 keV
	300 keV

	DD
	0.17
	0.17

	TT
	0.96
	0.65

	DT
	98.86
	99.18


3. Preparation of source definition card for use in simulations
3.1 Properties of source components
To quantify the details of different source components simulated using the ENEA-JSI source subroutine, each of the components was simulated separately. For each component 2×108 source events were recorded through MCNP’s PTRAC functionality and analyzed using a Python script. In PTRAC file MCNP outputs all relevant information about the particle history, e.g. its position, direction, energy, and statistical weight. Additionally, filters can be applied to reduce the size of the file. In our case only the neutron source events were written in the PTRAC file. The Python script (Python script 1 in Fig. 7) sorted these source events according to their emission angle relative to the ion beam, and into energy bins within each angular interval. 400 angular bins (equidistant in cosine scale) and 10 keV wide energy bins in relevant region were used to produce a high resolution description of the neutron source. The results of this process performed for each of the source components were 1D arrays describing the probabilities for emission in angular intervals, and 2D arrays describing neutron emission spectra for all emission angles.
When extracting the data, care needs to be taken that the energy and cosine boundaries do not cause rounding errors. As results in PTRAC file are shown in five digits, the cosines should be divided in intervals corresponding to the integer multiple of 10-6, and energies in integer multiple of 1 keV for neutron energies above 10 MeV.
3.2 Source definition card preparation
Once the relative intensities of different source components are determined the source description (e.g. SDEF card in MCNP) can be produced. The definition consists of the probability of emission into each of the angular bins, and of relative probabilities for emission at different energies, which correspond to 1D and 2D arrays from previous step, respectively.
To construct a single source description the 1D and 2D arrays describing different source components are multiplied with their respective relative intensity (Table 2 for DT components, and importance of DD and TT components in Table 4) and summed into combined 1D and 2D arrays (Python script 2 in Fig. 7). Once the source is described with a single pair of 1D and 2D arrays it is output into a source description file (SDEF). The process is visually described in Fig. 7.
A single SDEF source description can be used to describe all source components. However, if the objective is to determine the separate effects of different source components it is advisable to prepare separate source definitions for each of the components or at least for each of the reactions, i.e. DT, DD and TT separately.
[image: image9.png]do for all source components

Relative source
1ntensities

1-th PTRAC

i-th angular
emission intensity
(1D)

i-th neutron
spectra in angular
bins (2D)

SDEF. txt





Fig. 7. Schematic description of data extraction and MCNP source definition card (SDEF) preparation.

Due to low relative contributions of the DD and TT neutron source components they can be for most uses in compact neutron generator applications either neglected or simplified using isotropic spectra, i.e. their angular dependence neglected. This can be justified by previous analyses which show that the majority of the anisotropy is caused by the generator’s geometry and not by the anisotropy of fusion reactions [11].
4. Conclusions
The process of the neutron source spectrum analysis for a mixed beam DT neutron generator was proposed and demonstrated on a relevant example. Results show that in typical compact neutron generators with Eion of up to 110 keV the DD and TT neutron source components contribute less than 1 % to the total neutron emission. Our analysis found that this is still the case, for typical beam and target composition, even when the acceleration voltage is increased to 300 kV, as is typical for neutron generators with higher intensities. However, given the ratios of cross-sections, it is possible that contribution of components other than DT result in higher intensities under significantly different beam and target configurations. An example of relevant application for our analysis of source components is the use of an intense neutron generator for calibration of neutron detectors in future large fusion machines, e.g. ITER and DEMO, where DD and TT components will have to be taken into account if uncertainty in the neutron source will have to be significantly lower than 5 %. Additionally, the preparation process of the source definition card for use in simulations was explained.

Acknowledgments

This work has been carried out within the framework of the EUROfusion Consortium and has received funding from the Euratom research and training programme 2014-2018 under grant agreement No 633053. The views and opinions expressed herein do not necessarily reflect those of the European Commission.
The authors acknowledge the projects Development of methodology for calibration of neutron detectors with a 14.1 MeV neutron generator J2―6752, and Training of young researchers 1000-13-0106 were financially supported by the Slovenian Research Agency.
References
[1] P. Batistoni, et al., “Technical preparations for the in-vessel 14 MeV neutron calibration at JET,” Fus. Eng. Des., 117, pp. 107-117 (2017).
[2] V. Valjković, “14 MeV Neutrons: Physics and Applications”, CRC Press, Boca Raton, 2016.
[3] M.B. Chadwick, et al., "ENDF/B-VII.1: Nuclear Data for Science and Technology: Cross Sections, Covariances, Fission Product Yields and Decay Data," Nuclear Data Sheets, 112, (2011).
[4] A. Milocco, et al., “A Monte Carlo model for low energy D–D neutron generators,” NIMB, 271, pp. 6-12 (2012).
[5] T. Goorley, et al., “Initial MCNP6 Release Overview – MCNP6 version 1.0,” LA-UR-13-22934 (2013).

[6] A. Čufar, et al., “Comparison of DT neutron production codes MCUNED, ENEA-JSI source subroutine and DDT,” Fus. Eng. Des, 109-111, pp. 164-168 (2016).
[7] D. Rigamonti, et al., “Unprecedented very high resolution neutron spectroscopy measurements of 14 MeV neutrons and new prospects for diagnosing DT fusion plasmas,” will be submitted to New Journal of Physics, 2017.
[8] P. Batistoni, et al., 14 MeV Calibration of JET neutron detectors – Phase 1: calibration and characterization of the neutron source, submitted to Nuclear Fusion, (2017).
[9] M. Pillon, et al., “Characterization of a diamond detector to be used as neutron yield monitor during the in-vessel calibration of JET neutron detectors in preparation of the DT experiment,” Fus. Eng. Des., 106, pp. 93-98 (2016).
[10] P. Sauvan, et al., “New capabilities for Monte Carlo simulation of deuteron transport and secondary product generation,” NIMA, 614, pp. 323-330 (2010).

[11] A. Čufar, et al., “Calculations to support JET neutron yield calibration: Modelling of neutron emission from a compact DT neutron generator,” NIMA, 847, pp. 199-204 (2017).
_______________________________________________________________________________

author’s email: aljaz.cufar@ijs.si


1See the author list of “X. Litaudon et al 2017 Nucl. Fusion 57 102001″

_1562675219.unknown

_1562675238.unknown

