Two-phase flow measurements at high void fraction by a Venturi meter
Grazia Monni, Mario De Salve, Bruno Panella* 

Politecnico di Torino, Energy Department, Torino, 10129, Italy
*Corresponding Author, E-mail:bruno.panella@polito.it
KEYWORD: Venturi Flow Meter, pressure losses, two-phase flow rates, annular flow
Abstract
An experimental investigation of a vertical upward annular two-phase flow across a Venturi flow meter has been performed for the measurement of two-phase flow parameters with reference to the experimental simulation of nuclear accidents, as LOCA, characterized by very high void fraction. The pressure drops between the inlet and throat section and between inlet and outlet (irreversible pressure losses) have been measured and analyzed. The Venturi flow meter is characterized by an inlet diameter of 80 mm and a throat diameter of 40 mm (β=0.5), with equal convergent and divergent angles (θ=21°). The instrument has been tested in a test section having an internal diameter (Di) of 80 mm  and a total length (L) of about 4 m, with air-water two-phase flow at ambient pressure. The air superficial velocity ranged between 14 and 18 m/s while the water superficial velocity ranged between 0.0008 and 0.005 m/s, so that the flow pattern was annular and the corresponding  flow quality ranged between 0.78 and 0.96.  The dependence of the pressure drops on the phases velocities has been analyzed and modeled as a function of the superficial velocities of the phases, and the mass flow rate of the two phases has been derived from the present test data when the annular flow pattern does not change in a significant way. The flow quality can be evaluated with an accuracy of 5% and the  mass flow rate of air and water can be estimated with an accuracy better than 2% and 30% respectively. 
1. Introduction
The design of new nuclear reactors requires to carry out integral and separate effect tests on simulation facilities, as well as to perform safety systems verification and safety code validation. Within the framework of an Italian R&D program on Nuclear Fission, managed by ENEA and supported by the Ministry of Economic Development, the SPES3 (Pressurized Simulator for Safety Experiments) experimental facility[1], able to simulate the innovative small and medium size PWR (Pressurized Water Reactor) nuclear reactors, is being built and will be operated at the SIET Company laboratories in Piacenza (Italy). In such facility some design and beyond design basis accidents, like Loss of Coolant Accidents (LOCA), with and without the emergency heat removal systems, will be simulated. The experimental data, concerning a series of primary and secondary loops breaks, will be fundamental for the certification process of such reactors. The availability of such a complex plant allows us to perform other exploitation, and studies are foreseen for using it in a wider field of application for integral layout Small Modular Reactor (SMR) simulation. An accurate accident analysis requires the measurement of the mixture mass flow rate occurring in a LOCA, when a piping break occurs at high temperature and pressure. For this reason instruments and methodologies to evaluate mass flow rates, at the break and at other locations of the plant, need to be developed, considering the severe thermal-hydraulic conditions during the blow-down phase: the flow is critical, the average void fraction is higher than 90% and the flow pattern is annular or dispersed (wet-gas). 

Due to the simple design, the low costs, and the limited maintenance required, the Venturi flow meter (VFM) is one of the most used instruments for single-phase mass flow rate measurement, and it is one of the best candidate for two-phase mass flow rate measurement. In two-phase flow, the pressure drops are higher compared to single-phase flow, due to the strong interaction between the phases, that are strictly related with the phases distribution or flow pattern.  Considering its technical importance the response of the VFM in two-phase flow has been extensively studied by several authors [1]-[13], and a number of correlations and models have been developed. The well known works of Chisholm [9] and Murdock [10] have been used as guidelines for a new correlation developed by Bizon and Lin [14] and more recently by Moura and Marvillet [15], de Leeuw [16], Steven [17] and Xu [18]. All the correlations have been developed for very high or very low quality flows, and are focused on the correct measurement of the continuous phase, while the mass flow rate of the dispersed phase is usually neglected.  In some situations, as for example in nuclear safety simulation facilities, the measurement of the mass flow rate of the two phases is fundamental for a proper analysis of the accident evolution. So that the pressure drops in the VFM have to be modeled as a function of the two phases superficial velocities in order to evaluate the effect of each phase on the total flow meter pressure drops. In this context important information can be extracted from the  analysis of the irreversible pressure drops. Compared with the extensive works conducted on the analysis of the pressure drops between inlet and throat section, very few works have been performed on the modeling of the irreversible component.  Only recently new phenomenological model have been applied to analyze the pressure drops in those devices [19].

In the present research work, the experimental investigation of a vertical upward annular two-phase flow in a VFM has been performed. The dependence of the pressure drops on the phases velocities has been analyzed and modeled, and the differences with respect to the air single-phase flow have been evaluated.  

The predictions of the existing correlations are compared with the experimental pressure drops, and a formulation that predicts the experimental data is presented. Moreover the irreversible VFM pressure drops have been measured and analyzed. The experimental results have been modeled as a function of the flow quality and of the superficial velocities of the phases, showing that additional relevant information on two-phase flow can be extracted from the test data. In the last part of the work the characterization of the two-phase Venturi pressure drops is used to evaluate the mass flow rate of the two phases showing that further information can be derived from the measurement of the irreversible pressure drops. 

2. Venturi flow meter characteristics 

The derivation of the mass flow rate from the measurement of the differential pressure in a VFM is commonly used for single-phase flow in a wide range of industrial application. The technical standards [20] are, in this case, well defined and verified by several experimental and theoretical models.

The reference equation of the VFM in single-phase flow can be obtained applying the Bernoulli equation (balance between kinetic energy and static energy of a vena fluida) at the inlet and at the throat section: 


[image: image1.wmf](

)

2

1

2

2

2

1

2

v

v

p

p

p

-

×

=

-

=

D

r

      



(1)

so that the volumetric flow rate is given by: 


[image: image2.wmf](

)

Y

F

p

A

C

Q

a

d

×

×

ú

ú

û

ù

ê

ê

ë

é

÷

÷

ø

ö

ç

ç

è

æ

-

D

=

5

.

0

4

2

1

2

b

r



 
(2)
The correction factor Cd is introduced to take into account the vena contracta cross section and the additional frictional losses and viscosity and turbulence effects. The numerical values of Fa  are based on experimental data and are tabulated in ISO5167 [20]. For incompressible flow the coefficient Fa can be set equal to 1, while for gas and steam Fa>1 and it depends on the flow meter geometry and on flow conditions.

The Venturi (shown in Fig. 1) has been designed to operate in both fluid direction (symmetrical flow meter), with the angle of the convergent section equal to the divergent section angle (θ=21°), in order to be able to work in reverse flow conditions. 
[image: image3.jpg]



Fig. 1: Venturi flow meter picture
The selected material, Plexiglas, allows the direct flow visualization along the instrument. Compared with the Herschel model the present VFM is characterized by a higher divergent angle (usually set to 15°), with lower manufacturing costs but also higher irreversible pressure drops, due to the flow separation. The geometrical characteristics of the experimented VFM are described in Table 1.
	VFM characteristics

	D1
	80
	mm

	D2
	40
	mm

	β
	0.5
	-

	θconvergente = θdivergente
	21
	°

	Ltot_Venturi
	340
	mm

	Lupstream
	628
	mm

	Ldownstream
	628
	mm


Table 1: Venturi flow meter characteristics
The single-phase instrument behavior has been analyzed,  in order to obtain the VFM discharge coefficient Cd (equation 2) by using the experimental values of mass flow rate and pressure. Due to the low pressure drops inside the instrument, the air can be considered incompressible, and the parameters  Fa and Y have been assumed equal to one. The Reynolds number dependence can be approximated by an exponential law of type: 


[image: image4.wmf]b

d

a

C

Re

×

=








 (3)
with the parameters a and b (a=1.5054 and b=-0.0510) obtained from the experimental data (Fig. 2). 
[image: image5.emf]0.0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

0.9

1.0

2.E+04 4.E+04 6.E+04 8.E+04 1.E+05

Cd

Re


Fig. 2: Venturi discharge coefficient as a function of the Re number
The Re number is evaluated as:


[image: image6.wmf]m

r

1

Re

VD

=


 





(4)
In single-phase flow the discharge coefficient value takes into account the phenomena of the wall flow detachment at the inlet of the divergent section; in two-phase flow, the presence of the interface between the two phases complicates the physical interpretation of this correction coefficient. 
3. Venturi flow meter two-phase flow multiplier
According to the well known approach [2,6-10], the frictional pressure gradient of a two-phase mixture flowing in a pipe can be correlated to the single-phase pressure drop by means of the two-phase flow multiplier: 


[image: image7.wmf]2

g

g

TP

dz

dp

dz

dp

f

=

÷

ø

ö

ç

è

æ

-

÷

ø

ö

ç

è

æ

-

  





(5)
where –(dp/dz)g is the frictional gas pressure gradient obtained by using the actual air phase flow rate for two-phase flow . The two-phase flow multiplier is expressed as a function of the Martinelli parameter [2,3]:

[image: image8.wmf]g

l

dz

dp

dz

dp

÷

ø

ö

ç

è

æ

÷

ø

ö

ç

è

æ

=

2

c



 



(6)
defined as the ratio between the single-phase liquid and gas pressure gradients that are evaluated with the actual flow rates, Wl and Wg  for two-phase flow. 

Assuming that the two phases are flowing under turbulent regime and using the Blasius correlation for the single-phase friction factor coefficients calculation 
f = 0.0079/Re1/4 
         
 



(7)

the parameter 
[image: image9.wmf]2

c

  can be written as: 


[image: image10.wmf]125

.

0

5

.

0

875

.

0

2

1

÷

÷

ø

ö

ç

ç

è

æ

×

÷

÷

ø

ö

ç

ç

è

æ

×

÷

ø

ö

ç

è

æ

-

=

g

l

l

g

x

x

m

m

r

r

c


    

(8)
For a two-phase mixture flowing in orifice and Venturi devices, the parameter is usually written in the following form [8]: 


[image: image11.wmf]5

.

0

2

mod

1

÷

÷

ø

ö

ç

ç

è

æ

×

÷

ø

ö

ç

è

æ

-

=

l

g

x

x

r

r

c

  



 (9)
The two-phase pressure drops are then related to the two-phase flow multiplier, by means of empirical or semi-empirical correlations. 

Chisholm [9], for the calculation of the two-phase flow multiplier, proposed the correlation expressed as:


[image: image12.wmf](

)

2

mod

2

mod

1

c

c

f

+

+

=

C

C

g


 


 (10)

where 

[image: image13.wmf]5

.

0

5

.

0

1

÷

÷

ø

ö

ç

ç

è

æ

+

÷

÷

ø

ö

ç

ç

è

æ

=

l

g

g

l

S

S

C

r

r

r

r





 (11) 

and S is the slip ratio, that has been assumed equal to: 

[image: image14.wmf]4

/

1

÷

÷

ø

ö

ç

ç

è

æ

=

g

l

S

r

r





 

(12) 

Murdock [10], following the same approach, proposed: 


[image: image15.wmf](

)

(

)

2

mod

2

26

.

1

1

c

f

+

=

M

g






(13)
for an orifice,





[image: image16.wmf](

)

(

)

2

mod

2

5

1

c

f

+

=

M

g

   




(14)

for a Venturi.
4. Test section and experimental matrix 

The experimental facility consists of the feed water and the feed air loops, that are equipped with instruments to measure the single-phase flow parameters (flow meter, temperature and pressure), and of a vertical pipe (having a length of about 4 m and  an inner diameter of 80 mm), containing the test section. The water flow rate is changed by varying the pump rotation frequency and it is measured by a rotameter. The air flow is provided by a blower, that allows us to obtain high flow rate at low pressure (near ambient pressure), and it is measured by a calibrated orifice flow meter; whose accuracy is  2% full scale value. 

The test section, shown in Fig. 3, is transparent (Plexiglas) in order to visualize the flow pattern. Air enters axially into the test section and water flow is injected co-axially by means of a porous bronze; the mixing zone is located at 400 mm from the test section inlet. The 2.5 m long  test section is equipped with two pneumatic quick closing valves (QCV) to measure the volumetric void fraction; the uncertainty associated with the void fraction measurement has been estimated as ∆α = ±  0.0012. Downstream of the upper valve the two phases are separated in a tank at atmospheric pressure. 

In the test section an Electrical capacitance probe (ECP), having a total length of 1210 mm, used for void fraction measurement, is installed. The ECP results have been presented in ref. [21]. After the ECP the VFM is installed between two straight pipes of 1290 mm upstream and downstream respectively. The absolute pressure is measured at the inlet of the test section, while the VFM pressure drops are measured between the inlet and the throat sections and between the inlet and the outlet sections (irreversible pressure loss). The fluid temperature is measured at different location of the test section by means of thermocouples. 
[image: image17.png]— Air outlet

I

Water outlet

Water inlet —’ _.J

Airinlet  ——jp

VFM

ECP

acv

400 mm

175 mm

1290 mm

1210 mm

175 mm

400 mm





Fig. 3: Vertical Test Section schematic
The mixture void fraction is measured by means of the QCV technique. The instrument signals (single-phases parameters and signals from the test section instrumentation) are acquired by means of a NI DAQ, using the LabView® software. The acquisition time was equal to 30 s with an acquisition frequency of 1250 Hz.
The experiments have been carried out fixing the mass flow rate of the two phases at the inlet of the test section. The air superficial velocity ranged between 14 and 18 m/s while the water superficial velocity ranged between 0.0008 and 0.005 m/s, so that the flow pattern was annular and the corresponding void fraction is higher than 0.97, while the flow quality ranges from 0.78 to 0.96. 
For the tested superficial velocities range the variation of the void fraction (measured by QVC technique) is shown in Fig. 4. The small variation of this parameter is due to the very small values of the liquid superficial velocities.
In the present work the value of ρlJl  is much smaller than ρgJg. The liquid superficial velocity range has been chosen in order to obtain the annular flow pattern in the test section; considering the experimented air flow velocites, higher values of the liquid superficial velocities involve the presence of a counter-current flow.  
[image: image18.emf]14 15 16 17 18 19 20

0.96

0.965

0.97

0.975

0.98

0.985

0.99

0.995

1

J

g

 [m/s]



 

 

J

l

 = 0.0008 m/s

J

l

 = 0.0017 m/s

J

l

 = 0.0028 m/s

J

l

 = 0.0033 m/s

J

l

 = 0.0039 m/s

J

l

 = 0.0050 m/s


Fig. 4: Void fraction as a function of the superficial velocity of the two phases
5. Experimental results
5. 1
Venturi flow meter pressure drops 
In Fig.s 4 and 5 the experimental VFM pressure drops are plotted as a function of the flow conditions.  

In Fig. 5 the pressure drop dependence on the total mass flow rate is shown: the VFM pressure drop is a function of the mass flow rate of the two phases; introducing the ratio Wl/Wg, if this ratio is equal to zero the air single-phase flow is obtained, while increasing the fraction of water in the mixture an increase of the pressure drops is highlighted. The dependence on the superficial velocities of the two phases and the small effect due to the liquid flow rate is confirmed in Fig. 6. In Fig. 7 the two-phase pressure drop, normalized with reference to the gas single-phase pressure drop, is shown as a function of the phases superficial velocities: the two-phase flow value is very close to the gas single-phase flow value, with a maximum increase of about 7%.
[image: image19.emf]0.08 0.09 0.1 0.11 0.12 0.13

20

25

30

35

40

45

W

tot

 [kg/s]



p

V

 [mbar]

 

 

J

l

 = 0.0008 m/s

J

l

 = 0.0017 m/s

J

l

 = 0.0028 m/s

J

l

 = 0.0033 m/s

J

l

 = 0.0039 m/s

J

l

 = 0.0050 m/s


Fig. 5: VFM pressure drop as a function of the total mixture mass flow rate
The pressure drop increase in presence of two–phase flow is caused by the interaction between the gas and liquid phases: liquid droplets accelerated by the gas, irreversible drag force work done by the gas phase accelerating the liquid film and frictional wall losses, determine the magnitude of the observed pressure drop increase. The flow field is characterized and complicated by the continuous deposition and entrainment of liquid droplets along the Venturi length and by the presence of waves on the liquid film surface. The continuous deposition and entrainment process contributes to the overall pressure drop through the loss of momentum caused by the acceleration of the newly entrained droplets. 
[image: image20.emf]14 15 16 17 18 19

20

25

30

35

40

45

J

g

 [m/s]



p

V

 

-

T

P

 [mbar]

 

 

J

l

 = 0.0008 m/s

J

l

 = 0.0017 m/s

J

l

 = 0.0028 m/s

J

l

 = 0.0033 m/s

J

l

 = 0.0039 m/s

J

l

 = 0.0050 m/s


Fig. 6: VFM pressure drop as a function of the two phases superficial velocities
  [image: image21.emf]14 15 16 17 18 19

1

1.01

1.02

1.03

1.04

1.05

1.06

1.07

1.08

J

g

 [m/s]



p

V

-

T

P

 / 



p

V

-

g

 

 

J

l

 = 0.0008 m/s

J

l

 = 0.0017 m/s

J

l

 = 0.0028 m/s

J

l

 = 0.0033 m/s

J

l

 = 0.0039 m/s

J

l

 = 0.0050 m/s


Fig. 7: Normalized VFM pressure drop as a function of the two phases superficial velocities
The surface waves produce an effectively roughened surface over which the gas flows increasing the momentum loss due to the interfacial shear stress. On the ground of the previous considerations, the pressure drop in the VFM can be expressed as a function of the single-phase pressure drop with a correction factor (a two-phase flow multiplier) that can take into account the presence and the effect of the liquid phase at different flow qualities.  The two-phase flow multiplier can been written as: 


[image: image22.wmf]1

8

.

7

2

mod

,

+

×

=

c

f

g





          

(15)
where the two constants have been obtained as best-fit of the present experimental data. 

The two-phase flow multiplier values evaluated by the correlation (15) are compared with the correlations of Murdock, Chisholm and the present test data (Fig. 8): the classical correlations under estimate the two-phase multiplier 
[image: image23.wmf]2

g

f

.  

The experimental points dispersion can be justified by a change of the slip ratio value, that is considered constant in all the analyzed correlations.  In Fig. 9 the pressure drops evaluated with the new correlation are compared with the experimental values, showing a very good agreement and a calculation accuracy better than 5%. 

[image: image24.emf]0 0.002 0.004 0.006 0.008 0.01

1

1.02

1.04

1.06

1.08

1.1

1.12



exp



2g

 

 



2

M



2

C



2

mod



2

exp


Fig. 8: Two-phase flow multiplier: comparison between experimental data and correlations
[image: image25.emf]20 25 30 35 40 45

15

20

25

30

35

40

45



p

V

 experimental [mbar]



p

V

 predicted [mbar]

 

 

-5%

+5%


Fig. 9: VFM pressures drops: comparison between experimental data and predicted values (equation 15)
In the present work also the irreversible pressure losses, measured between the Venturi inlet and outlet sections have been analyzed and correlated with the flow parameters. In the following pictures the dependence of the VFM irreversible pressure loss on the flow parameters are highlighted. The analysis of this component, that usually are not considered in two phase flow models, allows us to understand better the effect of the dispersed phase in the two- phase pressure drops. 

In Fig. 10 the two VFM pressure drop components are shown as a function of the total mixture mass flow rate. 
Fig. 11 shows the irreversible pressure loss component with the pressure drop component measured between VFM inlet and throat sections. The circle points highlight the linear relation existing between inlet-outlet (irreversible component) and inlet-throat pressure drops in single-phase flow. The square points show that for an annular flow the relation between the two pressure components is not longer linear; the difference is clearly due to the liquid phase presence, so that the relation between the two components can be used to analyze the effect of the liquid flow rate and to estimate the liquid mass flow rate itself.
Fig. 12 shows the dependence of the irreversible pressure loss on the superficial velocities of the phases. In Fig. 13 the two-phase irreversible pressure loss is plotted as a function of the irreversible single-phase (air) component. While, due to the liquid phase presence, the inlet-throat pressure drop increases of about 10%, if compared to the single-phase flow, the irreversible pressure loss increases from about 20% to 100% depending on the liquid flow rate. 

[image: image26.emf]0.08 0.09 0.1 0.11 0.12 0.13

10

20

30

40

50

W

t

 [kg/s]



p

V

 [mbar]

 

 

in-throat

in-out


Fig. 10: VFM pressure drop and VFM irreversible pressure loss vs. total mass flow rate

[image: image27.emf]22 26 30 34 38 42

5

10

15

20

25



p

V

 [mbar]



p

i

r

r

 [mbar]

 

 

Two-phase

Single-phase (air) 

 
Fig. 11: VFM irreversible pressure loss vs. VFM pressure drops
[image: image28.emf]14 15 16 17 18 19

8

10

12

14

16

18

20

22

24

J

g

 [m/s]



p

V

i

r

r

 [mbar]

 

 

J

l

 = 0.0008 m/s

J

l

 = 0.0017 m/s

J

l

 = 0.0028 m/s

J

l

 = 0.0033 m/s

J

l

 = 0.0039 m/s

J

l

 = 0.0050 m/s


Fig. 12: VFM irreversible pressure loss as a function of the superficial velocities of air and water
The previous analysis allows the derivation of a correlation able to describe the irreversible pressure loss change as a function of the flow rate of the two phases. The proposed correlation, developed for the present tested conditions, expresses the irreversible pressure loss as a function of the gas superficial velocity and of the ratio between the liquid and the gas superficial velocities, highlighting the effect of the dispersed phase:


[image: image29.wmf](

)

(

)

4

1

3

2

k

J

J

J

k

p

k

g

l

k

g

g

irr

+

×

×

=

D

r




 (16)
where Jg and Jl are the superficial velocity of air and water respectively. The following constants have been obtained from the best fit of the experimental data:
k1 = 0.2096

k2 = 2

k3 = 0.13

k4= -2.9786

[image: image30.emf]6 8 10 12 14 16

6

8

10

12

14

16

18

20

22

24



p

irr-g

 [mbar]



p

i

r

r

 [mbar]

 

 

Two-phase

Single-phase (air)


Fig. 13: VFM pressure loss as a function of the air single-phase irreversible pressure loss

The comparison between the predicted irreversible pressure loss and the test data is shown in Fig. 14: the proposed correlation allows the estimation of the VFM irreversible pressure loss, in annular flow, at high void fraction, with an accuracy of 5%. 
[image: image31.emf]10 15 20 25

10

12

14

16

18

20

22

24

26



p

irr

 experimental [mbar]



p

i

r

r

 predicted [mbar]

+5%

-5%


Fig. 14: Comparison between predicted (equation 16) and experimental VFM irreversible pressure losses

The proposed correlation highlights the relationship between irreversible pressure losses and the two-phase flow parameters, Jg and Jl. The coefficients of correlation (16), obtained for the present experimental conditions of flow velocities, pressures and temperatures, need to be verified in a higher range of flow conditions in order to extend the validity of the correlation. 
5.2 Two-phase mass flow rate evaluation
The measurement of the two-phase flow mass flow rate is essential for several industrial applications. Typically a set of instruments (Spool Piece - SP) must be installed in order to evaluate the mass flow rate of the phases in a large range of flow patterns, pressures and temperatures [7,8]. Each instrument of the SP has to be sensitive to the different properties of the flow (like momentum, velocity, density, void fraction) and the selection of the instruments depends strongly on the experimental conditions: pressure, temperature and phases velocities. Different instruments can be coupled in a SP, and the VFM, characterized by the absence of moving parts and by a smoother flow profile than the orifice plates,  is one of the instruments more suitable to be installed in a wide range of industrial applications.
In single-phase flow the VFM allows the estimation of the mass flow rate from the pressure drop across the pipe restriction, but in presence of two phases the direct correlation between pressure drops and mass flow rate is not possible, so that additional information and a model able to interpret the signals are required [23]-[25]. 

In limited application ranges, if the flow pattern doesn’t change, as in the present experimental conditions, it is possible to extract the essential information from the signals of a VFM. 

The closure equation of the model in this case is based on the measurement of the irreversible pressure losses and on the proposed correlation (16). 

By means of an iterative approach (Fig. 15) the flow quality and the mass flow rate of the two phases have been estimated by using only the acquired signals of the VFM and the information concerning the absolute pressure and the temperature of the flow. During the first iteration the air flow rate is estimated considering a single-phase flow in the VFM (correlations (2) and (3)), while starting from the second iteration this value is corrected using the estimated value of the flow quality, by using the correlation (15): an initial guess value of this parameter is used to evaluate the VFM pressure drops, and the iterative loop goes on until the estimated pressure drop reaches the experimental value.  Then the mass flow rates of the two phases are estimated by using the flow quality value and the VFM pressure losses correlation (16). The reference signals have been obtained as the mean values of the 30 s acquisition time. 
[image: image32.emf]5.0

2

1



































l

g

x

x







Δp

V

x -W

TP

T, P

Δp

irr

g l

g l

 

 





 

 

4 1

3

2

k J J J k p

k

g l

k

g g irr TP

    





x

guess

V g V TP

p p

 

     

) 1 8 . 7 (



W

g

=x·W

TP

W

l

=(1-x)·W

TP


Fig. 15: Mass flow rate estimation: iterative procedure flow chart

In Fig. 16 the estimated flow quality values are compared with the experimental data, obtained from the measurement of the single-phase mass flow rates. The relative error is in each case lower than 5%.
[image: image33.emf]0.75 0.8 0.85 0.9 0.95 1

0.75

0.8

0.85

0.9

0.95

1

x

exp

x 

e

s

t

+5%

-5%


Fig. 16: Comparison between experimental and estimated flow quality
In Fig.s 17 and 18 the estimated mass flow rates of air and water are compared with the measured values: the air mass flow rate is estimated with an accuracy of 2%, while a lower accuracy characterizes the liquid flow rate prediction (the relative error is lower than 30% for all conditions).
[image: image34.emf]0.08 0.085 0.09 0.095 0.1 0.105 0.11

0.08

0.085

0.09

0.095

0.1

0.105

0.11

W

g-exp

 [kg/s]

W

g

-

e

s

t

 [kg/s]

-2%

+2%


Fig. 17: Comparison between experimental and estimated air single-phase mass flow rates
[image: image35.emf]0 0.005 0.01 0.015 0.02 0.025 0.03 0.035

0

0.005

0.01

0.015

0.02

0.025

0.03

0.035

W

l-exp

 [kg/s]

W

l

-

e

s

t

 [kg/s]

+30%

+20%

-20%

-30%


Fig. 18: Comparison between experimental and estimated water single-phase mass flow rates
6. Conclusions 
In the present research work the experimental investigation of a vertical upward annular two-phase flow in a Venturi Flow Meter (VFM), with reference to the experimental simulation of nuclear accidents, as LOCA, characterized by very high void fraction.has been performed. The dependence of the pressure drops, evaluated between  the VFM inlet and throat and between the inlet and outlet sections (irreversible pressure loss), on the characteristic flow parameters (flow velocities, quality and void fraction) have been analyzed and discussed. 

Correlations describing the relation between velocities and VFM pressure drops have been proposed for the two pressure drops components. For both correlations, the prediction accuracy is better than 5%.  

The paper highlights that, from the measurement of the VFM irreversible pressure losses, important information can be extracted concerning the effect of the liquid dispersed phase on the total pressure drops. Due to the liquid phase presence, the inlet-throat pressure drop increases of about 10% compared to the single-phase flow, while the irreversible pressure losses increase from about 20% to 100% depending on the liquid flow rate. The proposed correlation describes the irreversible pressure loss as a function of the flow rate of the two phases, highlighting the effect of the dispersed phase. 

In addition the mass flow rate of the two phases and the flow quality have been evaluated by analysing  the test data for the present conditions of a well established annual flow pattern with high void fraction. The developed calculation procedure allows the evaluation of the flow quality of the mixture with an accuracy of about  5% and the estimation of the mass flow rate of air and water with an accuracy better than 2% and 30% respectively. 
Nomenclature
	A2
	[m2]
	throat cross section

	Cd
	[-]
	discharge  coefficient  

	D1
	[m]
	inlet section diameter

	D2
	[m]
	throat section diameter

	f
	[-]
	friction factor

	Fa
	[m/s]
	correction factor for thermal expansion

	J
	[m/s]
	superficial velocity

	L
	[m]
	geometrical length

	p1 - p2              
	[Pa]
	from inlet to throat section pressure drops

	Q
	[m3/s]
	volumetric flow rate 

	Re
	[-]
	Reynolds number

	S
	[-]
	slip ratio

	v1 
	[m/s]
	inlet section velocity

	v2
	[m/s]
	throat section velocity

	W
	[kg/s]
	mass flow rate

	x
	[-]
	flow quality

	Y
	[-]
	compressibility factor

	Greek symbols

	α
	[-]
	void fraction

	β
	[-]
	diameters ratio   D2/D1

	ρ
	[kg/m3]
	fluid density

	θ
	[°]
	divergent / convergent angle

	µ
	[Pa·s]
	dynamic viscosity

	Φ
	[-]
	two-phase flow multiplier

	χ
	[-]
	Martinelli parameter

	∆p
	[Pa]
	pressure drop

	Subscripts

	d
	 
	discharge

	exp
	 
	experimental 

	est
	 
	estimated

	g
	 
	gas

	int
	 
	internal

	irr
	 
	irreversible

	l
	 
	liquid

	tot
	 
	total (liquid + gas)

	TP
	 
	two-phase

	V
	 
	Venturi 


Abbreviations
ECP
Electrical capacitance probe
QCV
Quick closing valves

VFM

Venturi flow meter

Ackowledgement

The present research has been supported by ENEA and by the Ministry of Economic Development.

The authors thank R. Costantino and G. Vannelli for their support in the laboratory. 

References
[1] R. Ferri, A. Achilli, C. Congiu, et al., SPES3 facility and IRIS reactor numerical simulations for the SPES3 final design, Proc. European Nuclear Conference (ENC '10), Barcelona, Spain, May- June 2010.
[2] R. W. Lockhart and R.C. Martinelli, Proposed correlation of data for isothermal two phase two-component flow in pipes,  Chem. Eng. Prog., 45, 39 (1949).
[3] R. C. Martinelli and D. B. Nelson, Prediction of pressure drop during forced circulation boiling of water , Trans. ASME, 70, 695 (1948). 
[4] A. E. Dukler, M. Wicks, and  R.G. Cleveland, Frictional pressure drop and holdup in two phase flow, Part A, A  Comparison of existing correlations for  pressure  drop  and  holdup,  B  An approach through similarity  analysis,  AIChE  J.  10, 38-43  and 44 -51 (1964).
[5] K. Aziz, G. W. Govier, M. Fogarasi, Pressure drop in wells producing oil and gas, J. Canadian Petro. Tech, 38-48  (1972).
[6] C. J. Baroczy,  A systematic correlation for two phase pressure drop,  AIChE reprint 37 presented at 8th National Heat Transfer Conference, Los Angeles, August (1965).
[7] M. De Salve, G. Monni, B. Panella, State of art and Selection of techniques in Multiphase Flow Measurement, Report RdS/2010/67, ENEA (2011).
[8] C. R. Baker, Flow measurement handbook, Cambridge University Press (2000).
[9] D. Chisholm, Two phase flow through sharp edged orifices. J. Mech. Eng. Sci., 19(3), 128–30 (1977).

[10] J. W. Murdock, Two phase Flow Measurement with Orifices, ASME J. Basic Eng., December, 419-433 (1962).

[11] W. Jitschin, Gas flow measurement by the thin orifice and the classical Venturi tube, Vacuum 76 (2004).
[12] J. R.  Fincke, Performance Characteristics of an Extended Throat Flow Nozzle for the Measurement of High Void Fraction Multi-Phase Flows, Lockheed Martin Idaho Technologies (1999).
[13] L.D. Fang, T. Zhang, N.D. Jin, A comparison of correlations used for Venturi wet gas metering in oil and gas industry. J. Petroleum Sc. and Eng. (2007).
[14] Z.H. Lin, Two phase flow measurements with sharp-edged orifices. Int. J. Multiphase Flow (1982). 

[15] L.F.M., Moura, C. Marvillet, Measurement of Two phase Mass Flow Rate and Quality Using Venturi and Void Fraction Meters. Proc. 1997 ASME International Mechanical Engineering Congress and Exposition, Dallas, TX, USA (special issue), Fluids Eng. Division, FED 244 (1997).
[16] H. De Leeuw, Venturi meter performances in wet gas flow. In: B.H.R. group multiphase (1997).

[17] R. Steven, Wet gas metering. Ph.D. thesis. Glasgow: University of Strathclyde (2001).

[18] L. Xu, X. Jian, D. Feng and T. Zhang  , On fluctuation of the dynamic differential pressure signal of Venturi meter for wet gas metering, Flow Meas. and Instr.,14 (2003).
[19] A.H.A. M. Hasan, G.P. Lucas, Experimental and theoretical study of the gas-water two phase flow through a conductance multiphase Venturi meter in vertical annular (wet gas) flow, Nuclear Eng. and Design, 241, 1998-2005 (2011).
[20] Technical guideline ISO 5167-1:4 (E) (2003). 
[21] G. Monni, M. De Salve, B. Panella, and C. Randaccio, Electrical Capacitance Probe Characterization in Vertical Annular Two phase Flow, Science and Technology of Nuclear Installations, Vol. 2013, Article ID 568287, 12 pages, ISSN: 1687-6075 (2013).
[22] M. Ishii, K. Mishima, Two-fluid model and hydrodynamic constitutive relations, Nuclear Eng. and Design,82, 107-126 (1984).
[23] H.J. Zhang, W.T. Yue, Z.Y. Huang, Investigation of oil–air two phase mass flow rate measurement using venturi and void fraction sensor, J. Zhejiang University Science 6A (6) (2005).
[24] Z.Z. Meng, Z.Y. Huang, H.F. Ji, H.Q. Li, Y. Yan, Air-water two phase flow measurement using a Venturi meter and an electrical resistance tomography sensor, Flow Meas. and Instr., 21, 268-276 (2010). 
[25] M. De Salve, G. Monni, B. Panella, A Model for a Spool Piece made up of Venturi and Void fraction Flow Meter in Horizontal Flow, Trans. ANS, Vol. 108 1013-1016, Atlanta, Georgia, June 16–20 (2013).
Disclosure policy

The authors declare that there is no conflict of interests regarding the publication of this article.

_1453561479.unknown

_1453562242.unknown

_1453879146.unknown

_1453879164.unknown

_1453613984.unknown

_1453562234.unknown

_1452331190.unknown

_1453016909.unknown

_1453016916.unknown

_1453016902.unknown

_1452331214.unknown

_1451735444.unknown

_1451736331.unknown

_1451994025.unknown

_1451995599.unknown

_1451736276.unknown

_1451538721.unknown

